## Cover Sheet
SPOT / NPEWS User Interface Specification V1.0
This document is the baseline user interface specification for SPOT NPEWS systems. This has been created based on solutions previously procured by Trusts and feedback from clinicians and technical colleagues.
This cover sheet will outline some of the guiding principles for the use of this document by ICSs and Trusts.
Use of MOSCOW
We have used the MOSCOW rating system to assess the priority of each requirement. In an ideal situation a solution would meet or exceed every requirement but we recognise that this isn't practical and have used the rating method to reflect differing priorities. Different people may use the MOSCOW ratings in different ways so we have provided a definition of our usage to make things clearer for users of this document
Must - Any SPOT / NPEWS solution must meet all of these requirements. When undergoing a procurement exercise no solution should be selected that fails one or more Musts.
Should - Ideally a SPOT / NPEWS solution would have all the Shoulds as well as all the musts. There may be some shoulds that don't apply to certain solutions, or some suppliers may do the majority of the Shoulds but miss one or two. Missing a should is not a barrier to a successful procurement but would be a useful comparison to make between competing suppliers to pick the one best suited.
Could - A could is a nice to have, these are often future looking or ideal features that it is understood many or most suppliers won't deliver yet, but may have on roadmaps. A lack of coulds should not be a barrier to supplier selection, but where several suppliers are similar in their ability to meet Musts and Shoulds whether they meet Coulds is a useful measure for differentiation.
Won't - This is used to call out something we explicitly wont be doing. No solution should do this feature as to do so would be an error. We include some won'ts to avoid assumptions that might be made otherwise
Use with other documents
This user interface (UI) specification is designed to complement the main technical specification for SPOT-NPEWS. These should be used together. A feature missing from this specification just means that we don't have specific UI concerns around the feature. It should still be implemented as per the technical specification but suppliers are free to make decisions on how best to implement it. To go with this document we have produced wireframes. These wireframes are designed to explain some of the concepts in this specification. The wireframes aren't designed to be a strict interpretation but to illustrate points raised.
Usage of this document by Integrated Care Systems (ICSs) and Trusts
The purpose of this document is to form a baseline for the SPOT / NPEWS solution for any ICS or trust who needs to procure a SPOT / NPEWS solution. This document is intended to be the national baseline, so as such it would be expected that any solution procured must use all of the Must requirements that are included. Given that this is a national document it is understood that we will be unable to reflect any locally specific requirements. As such it is expected that the trust add these to the requirements list before going out to procurement. For example we can add generic statements about integration, but the Trust may want to add a list of systems that they have that the solution must integrate with. There may also be some of the Shoulds or Coulds that change locally, for instance we might have included a could that just isn't relevant in your local environment so you can take it out. Alternatively one of our Shoulds may be a Must in your local area, so you should feel free to upgrade it's status. This is a baseline so we are more than happy for Trusts to increase the classification of items and aim at a higher standard. In terms of downgrades, care must be given before downgrading a Must, as you may find that your system isn't compliant with national standards, however there may be limited situations where you need to replace a generic must with a more specific one reflecting local realities.
Requirements from this document are being shared with suppliers in the SPOT / NPEWS market for comment and review so using this baseline should assist your procurements as suppliers will already have sight of it. It is hoped that by including shoulds and coulds into this specification we can help to inform suppliers roadmaps to resolve those issues where they might be diverging from the expected standard.
Key definitions
The Solution
When we use the term "the solution" (typically in the format of "The solution must do X") we mean the full set of system or systems provided by the supplier. No judgement is made as to whether this needs to be one piece of software or several that are integrated. The integration is a key assumption that is made, it wouldn't be acceptable to have a solution that had bits of technology that didn't speak to each other so clinicians were required to double enter data. However as long as it's fully integrated any number of software/hardware items can be included as part of the solution. It can be worth reading this as meaning "As a whole the fully integrated solution Must do X", but we have left that wording out for the sake of space. In terms of any requirements that refer to compliance (e.g. GDPR) this refers to each and every component part of the solution. So for a supplier to confirm that their solution meets this requirement they are confirming that all the related software they'll be deploying meets that requirement.
MOSCOW within requirements
When writing a requirement the relevant MOSCOW term will be used throughout the requirement (e.g. The solution should do X and it should also do Y and it should output in the format of z). Even though the less mandatory terms of Should and Could are used in this way, if a supplier is answering that they meet the requirement they are saying that they meet all of it. If they are in a situation where they meet some of the requirement but not all of it they need to make this clear in their submission. This can then be taken into account by the trust when assessing suppliers. For shoulds and coulds it isn't an automatic barrier if a supplier is only partially compliant, but if they just say that they are compliant and don't outline the areas that they are deficient then this could become an issue with any contract signed based on the supplier appearing fully compliant with a requirement.
Clinicians
Clinicians is our generally used term to refer to staff working with the SPOT / NPEWS system. For some requirements specific jobs are called out but otherwise clinicians is used as a generic term. Occasionally users is used as well, this means all users of the system both clinicians and others (e.g. technical administration staff). When clinicians is used it should not be taken to imply any particular job role, if the supplier needs clarification of which are or are not included locally they should seek this from the trust, as in some areas this will change depending on local needs.
ICS/Trust
The Integrated Care Systems (ICSs) and Trusts are used interchangeably through this document. Depending on the area purchasing of systems may be done at the ICS or at the Trust level and so both are used throughout. If the procurement is being done at ICS level requirements that refer to the Trust should be seen to refer to the ICS and visa versa.
Standards and datasets
When standards and datasets are listed it is assumed that the current version as of the date of procurement publication is to be used. Version numbers have not be used in this document to avoid creating an overhead in terms of updating them, but trusts are welcome to include them in their procurement documents. Any new datasets will be included as part of the regular updates to this document.
Patient
We've used patient to refer to the child or young person accessing medical services and having observations taken by clinicians
Carer
We've used carer to refer to the carer or parent responsible for the child or young person receiving medical care
Form
We've used Form generically to describe the interface used to input observation data on the PEWS system, in a similar way to the paper version being called a form. We understand that different supplier solutions may use different terms like interface, app or workflow, and in their systems "form" may mean something different. For our purposes we are using the term in it's generic plain English form and do not preclude different ways that suppliers may implement the described interface in their systems
Compliance with the standard
For a trust or a supplier to claim that they are compliant with the National PEWS standard in communications or advertising their solution must meet all of the must requirements in the specifications. A failure to meet any must have requirement will mean that the trust or supplier will not be allowed to claim to be compliant with PEWS.

## UI Requirements
General Requirements
ID | Requirement | MOSCOW
U1.1 | The solution must be able to support a colour blindness friendly mode for clinicians that have issues seeing colour. This must still make it clear which level applies so that the PEWS form can still be easily interpreted. | Must
Form Creation and Selection Requirements
ID | Requirement | MOSCOW
U2.1 | No requirements in this section
Viewing and Recording Observations Requirements
ID | Requirement | MOSCOW
U3.1 | The solution must allow the clinician to zoom in and out when viewing charts to see them in more or less detail as needed. This is to let them see fewer observations in more detail, or more observations in less detail. | Must
U3.2 | The solution must default to a clear view that shows recent trends. The amount of results shown may need to vary depending on how many observations have been recently conducted | Must
U3.3 | The solution should have some quick view options to change what is shown without manual scrolling by the clinician. These should include "Today" "This Week" "This Month". | Should
U3.4 | The solution should have the option to show the values of the current observations to the clinician without charting / trends. This might be an alternative screen, but should allow a clinician to see most recent values for each of the observations taken (without the previous values and trending) | Should
U3.5 | The solution must always be able to show the exact value of the observation to the viewing clinician | Must
U3.6 | The solution should show the exact value of the observation in an easy to see way. This should be either on a separate chart line for "value", or next to the dot, or on an interaction (like click/press/hover) | Should
U3.7 | The solution must arrange the y axis values from highest numeric values at the top to lowest numeric value at the bottom for each observation measure. For those measures without numeric values the solution must show these from most severe/highest scoring to least severe/lowest scoring | Must
U3.8 | The solution could present a visual indicator that the patient has been moved to theatre, scans or other care locations to explain why observations have paused | Could
U3.9 | The solution must make it obvious when any observation is skipped when a clinician is reviewing observations | Must
U3.10 | The solution won't draw a line over a skipped observation when graphing. Any skipped observation must cause a break in the line. If a set of observations is conducted but one is skipped (e.g. Blood Pressure) no line should be dawn over the two closest values, as this would imply a trend that might not be the case. The wireframes have an example of the break in the line. | Wont
U3.11 | The solution must ensure the clinician can always see the date/time of the observation(s) they're looking at so they always know which one(s) they're reviewing. | Must
U3.12 | The solution should have a jump to present option to quickly allow the clinician to return to the present set of observations | Should
U3.13 | The solution must distinguish between trend lines and observations so it is clear where an observation has been taken. Typically these are plotted as a dot, but any appropriate symbology is acceptable. The wireframes give examples of how this can be seen. | Must
U3.14 | The solution must plot the observation exactly in the box, this means that even if two values would fall in the same box if one is higher than the other this must be shown. For example both 41 and 49 would fall in the 40-50 box, but one should be at the bottom of the box and one at the top. | Must
Correcting and Deleting Observations Requirements
ID | Requirement | MOSCOW
U4.1 | No requirements in this section
Respiratory Rate Requirements
ID | Requirement | MOSCOW
U5.1 | The solution must allow a numeric value to be input for Respiratory Rate | Must
U5.2 | The solution could allow a clinician to input Respiratory Rate via an easy to use input method, e.g. a scroll, a value picker, auto dication or a number pad | Could
U5.3 | The solution must limit values for Respiratory Rate to exact numbers (integers) | Must
U5.4 | The solution must plot the exact point on the chart for the Respiratory Rate and show this via a marker on the chart that this is an observation not a trend, for example by plotting a dot | Must
U5.5 | The solution must plot a line between observations of Respiratory Rate to create a graph | Must
U5.6 | The solution must be able to show the reason an observation was skipped when requested by the clinician for Respiratory Rate | Must
U5.7 | The solution could limit the values of respiratory rate to possible results. The lowest possible value could be 0 and the highest possible value could be 150. | Could
Respiratory Distress Requirements
ID | Requirement | MOSCOW
U6.1 | The solution must present the different options for Respiratory Distress and allow the clinician to pick from them | Must
U6.2 | The solution should offer visible distinction (e.g. colour or sections) to show which types of Respiratory Distress will trigger which category | Should
U6.3 | The solution must either present the selected category of respiratory distress in the box and show the colour graded, or present a scale of categories of respiratory distress and mark which was selected. | Must
U6.4 | The supplier must be able to adapt to either respiratory distress option if one or the other is selected as the standard in future by NHS England | Must
U6.5 | The solution must allow the clinician to see the exact types of respiratory distress observed by interacting with the recorded category | Must
U6.6 | The solution must allow the clinician to easily record that no respiratory distress was observed. This must be distinct from any option to skip the observation. This must be the first option at the top of any input interface | Must
U6.7 | The solution must present the selection options for respiratory distress from None to Mild to Moderate to Severe so that it is easy for the clinician to find what they're looking for | Must
U6.8 | The solution must be able to show the reason an observation was skipped when requested by the clinician for Respiratory Distress | Must
Oxygen Saturation Requirements
ID | Requirement | MOSCOW
U7.1 | The solution must allow a numeric value to be input for Oxygen Saturation | Must
U7.2 | The solution could allow a clinician to input Oxygen Saturation via an easy to use input method, e.g. a scroll, a value picker, auto dication or a number pad | Could
U7.3 | The solution must present oxygen saturation in one of the following options: show the recorded value of oxygen saturation in the box and show the colour graded, OR present a scale of ranges of oxygen saturation and mark which was selected OR present a graph of oxygen saturation and plot points and graph it | Must
U7.4 | The solution could have an option to expand and contract between the full graph and the range based options | Could
U7.5 | The supplier must be able to adapt to any one of the oxygen saturation options if one of them is selected as the standard in future by NHS England | Must
U7.6 | The solution must limit values for Oxygen Saturation to exact percentages (integers) | Must
U7.7 | The solution must be able to show the reason an observation was skipped when requested by the clinician for Oxygen Saturation | Must
U7.8 | The solution could limit the values of oxygen saturation to possible results. The lowest possible value must be 20% and the highest possible value must be 100%. | Could
Respiratory Support Device Requirements
ID | Requirement | MOSCOW
U8.1 | The solution must present the different options for Respiratory Support and allow the clinician to pick from them | Must
U8.2 | The solution should offer visible distinction (e.g. colour or sections) to show which types of Respiratory Support will trigger which category | Should
U8.3 | The solution must present the category of respiratory support in the box and show the colour graded. | Must
U8.4 | The solution must allow the clinician to see the exact types of respiratory support recorded by interacting with the recorded category | Must
U8.5 | The solution must allow the clinician to easily record that no respiratory support was being offered. This must be distinct from any option to skip the observation. This must be the last option at the bottom of any input interface | Must
U8.6 | The solution must be able to show the reason an observation was skipped when requested by the clinician for Respiratory Support | Must
Oxygen Delivery Requirements
ID | Requirement | MOSCOW
U9.1 | The solution must allow a numeric value to be input for Oxygen Delivery. It must allow this to be either a percentage or a litres per minute value and identified as such | Must
U9.2 | The solution could allow a clinician to input Oxygen Delivery via an easy to use input method, e.g. a scroll, a value picker, auto dication or a number pad | Could
U9.3 | The solution must limit inputs for Oxygen Delivery using a litres per minute measure to two decimal places | Must
U9.4 | The solution should limit values for Oxygen Delivery using a oxygen percentage measure to exact numbers (integers) | Should
U9.5 | The solution must plot the exact point on the chart for the Oxygen Delivery and show this via a marker on the chart that this is an observation not a trend, for example by plotting a dot | Must
U9.6 | The solution must plot a line between observations of Oxygen Delivery to create a graph | Must
U9.7 | The solution must recognise when the measure for Oxygen Delivery changes between percentage and litres per minute and create a break in the line plotted on the graph | Must
U9.8 | The solution must be able to show the reason an observation was skipped when requested by the clinician for Oxygen Delivery | Must
U9.9 | The solution could limit the values of oxygen delivery to possible results. The lowest possible value when using percentages could be 16% and the highest possible value could be 100%. The lowest possible value when using litres per minute could be 0.02 and the highest possible value could be 16. | Could
Heart Rate Requirements
ID | Requirement | MOSCOW
U10.1 | The solution must allow a numeric value to be input for Heart Rate | Must
U10.2 | The solution could allow a clinician to input Heart Rate via an easy to use input method, e.g. a scroll, a value picker, auto dication or a number pad | Could
U10.3 | The solution should limit values for Heart Rate to exact numbers (integers) | Should
U10.4 | The solution must plot the exact point on the chart for the Heart Rate and show this via a marker on the chart that this is an observation not a trend, for example by plotting a dot | Must
U10.5 | The solution must plot a line between observations of Heart Rate to create a graph | Must
U10.6 | The solution must be able to show the reason an observation was skipped when requested by the clinician for Heart Rate | Must
U10.7 | The solution could limit the values of heart rate to possible results. The lowest possible value could be 0 and the highest possible value could be 300. | Could
Blood Pressure Requirements
ID | Requirement | MOSCOW
U11.1 | The solution must allow a numeric value to be input for both Systolic and diastolic Blood Pressure | Must
U11.2 | The solution could allow a clinician to input Blood Pressure via an easy to use input method, e.g. a scroll, a value picker, auto dication or a number pad | Could
U11.3 | The solution should allow a numeric value to be input for the Mean Blood Pressure | Should
U11.4 | The solution should limit values for Blood pressure to exact numbers (integers) | Should
U11.5 | The solution must plot the exact point on the chart for the systolic and diastolic Blood Pressure. This must use either the standard outward pointing arrow marking or the standard inward pointing arrow marking and have a line joining the systolic and diastolic values. | Must
U11.6 | The solution should plot the mean Blood Pressure with a dot on the line if a mean blood pressure is input | Should
U11.7 | The solution wont plot a horizontal graph line between different observations of systolic and diastolic Blood Pressure | Wont
U11.8 | The solution could plot a line between the mean Blood Pressure values to show trends | Could
U11.9 | The solution must be able to show the reason an observation was skipped when requested by the clinician for Blood Pressure | Must
U11.10 | The solution could limit the values of blood pressure to possible results. The lowest possible systolic value could be 0 and the highest possible value could be 400. The lowest possible diastolic value could be 0 and the highest possible value could be 400. | Could
Capillary Refill Time Requirements
ID | Requirement | MOSCOW
U12.1 | The solution must allow a numeric value to be input for Capillary Refill time | Must
U12.2 | The solution could allow a clinician to input Capillary Refill time via an easy to use input method, e.g. a scroll, a value picker, auto dication or a number pad | Could
U12.3 | The solution should limit values for Capillary Refill time to whole numbers (integers) | Should
U12.4 | The solution must either present the recorded value of Capillary Refill time in the box and show the colour graded, or present a scale of ranges of Capillary Refill time and mark which was selected. | Must
U12.5 | The supplier must be able to adapt to either Capillary Refill time option if one or the other is selected as the standard in future by NHS England | Must
U12.6 | The solution must be able to show the reason an observation was skipped when requested by the clinician for Capillary Refill time | Must
U12.7 | The solution could limit the values of Capillary Refill time to possible results. The lowest possible value could be 1 and the highest possible value could be 10. | Could
Carer Question Requirements
ID | Requirement | MOSCOW
U13.1 | The solution must allow a clinician to pick which escalation to trigger in response to a worse response to the carer question | Must
U13.2 | The solution must allow a clinician to record if the carer was asleep or unavailable. This must be distinct from skipping the observation | Must
Disability and Exposure Requirements
ID | Requirement | MOSCOW
U14.1 | The solution should remind the user to complete a Glasgow Coma Score sheet when AVPU indicates it. This may either be a text prompt or if a digital solution exists in that trust a direct link to it. | Should
U14.2 | The solution should allow the clinician to input a numerical value or pick from a visual scale for the pain score | Should
U14.3 | The solution could allow a clinician to input the pain score via an easy to use input method, e.g. a scroll, a value picker, auto dication or a number pad | Could
U14.4 | The solution should allow a trust to define the range for pain scores and present this to clinicians accordingly. This could be a numeric value or a locally defined set of images | Should
U14.5 | The solution must allow a new instance of sepsis to be recorded on it's own line, but also trigger the specific concern process if sepsis is recorded | Must
Temperature Requirements
ID | Requirement | MOSCOW
U15.1 | The solution must allow a numeric value to be input for Temperature | Must
U15.2 | The solution could allow a clinician to input Temperature via an easy to use input method, e.g. a scroll, a value picker, auto dication or a number pad | Could
U15.3 | The solution must limit values for Temperature to numbers with one decimal place | Must
U15.4 | The solution must plot the exact point on the chart for the Temperature and show this via a marker on the chart that this is an observation not a trend, for example by plotting a dot | Must
U15.5 | The solution must plot a line between observations of Temperature to create a graph | Must
U15.6 | The solution must be able to show the reason an observation was skipped when requested by the clinician for Temperature | Must
U15.7 | The solution could limit the values of temperature to possible results. The lowest possible value could be 0 and the highest possible value could be 45. | Could
Clinical Intuition Requirements
ID | Requirement | MOSCOW
U16.1 | The solution must allow a clinician to pick which escalation to trigger in response to a "Yes" response to the clinical intuition question | Must
U16.2 | The solution must allow the clinician to skip the clinician intuition question if they wish to | Must
U16.3 | The solution could treat an "I'm not sure" option the same as a skip but record it differently for reporting purposes | Could
Scoring and Escalation Requirements
ID | Requirement | MOSCOW
U17.1 | The solution must allow a clinician to pick which escalation to trigger in response to a specific concern being logged | Must
U17.2 | The solution must allow the clinician to break out of conducting observations at any point to trigger an escalation | Must
U17.3 | The solution must present the option to break out of observations to escalate in an obvious manner that is easy for a clinician to spot | Must
U17.4 | The solution must allow the clinician to select the reason they are escalating (e.g. Clinical intuition, Carer concern, specific concern) and then pick the escalation level. This must default to the highest level triggered by inputs to the various questions and observations but allow the clinician to override | Must
U17.5 | The solution must allow the observations to be continued after the patient has been escalated if needed | Must
U17.6 | The solution must present an alert at the end of the observation process letting the clinician know what level the patient has come out as, and what the guidance is for that level | Must
U17.7 | The solution should present a visually distinct warning if the escalation category increases (for whatever reason; PEWS, Specific Concern, Carer Concern, Clinical Intuition) rather than just remaining at the same level (e.g. moved from Medium to High, rather than staying at Medium). | Should
U17.8 | The solution must present a visual indicator that the patient has been escalated during/after a specific set of observations in addition to completing those sections of the form | Must
U17.9 | The solution must always have the option to present escalation guidance whenever an escalation is being offered to the clinician regardless of where they are in the process | Must
Comments Requirements
ID | Requirement | MOSCOW
U18.1 | No requirements in this section
Alarms Parameters and Risks Requirements
ID | Requirement | MOSCOW
U19.1 | No requirements in this section
Guidance Requirements
ID | Requirement | MOSCOW
U20.1 | The solution must show escalation guidance using the same colour codes as the form uses (with Red for emergency) so that it is clear to users what each escalation corresponds to. | Must